## 3.1 Ceramic
NASA 4602
Laboratory Design Criteria
Group | EM 40 |  |  |  | Space Name |  | Ceramic Matrix Composites Processing Laboratory
Exist. Loc. | 4612 / 1113 & 1117 partial |  | Date: | 2007-03-26 | Program Number
General |  |  |  |  | Plumbing
Unit Area | 845 |  |  |  | Sinks | 25x15x12 | Acid Waste Sys
Qty. | 1 |  |  |  | Cold Water | at sink | Hot Water | at sink
Hrs Occupancy | 10 |  |  |  | Pure Water |  | Grade
# Occupants |  |  |  |  | Comp Air |  | Pressure/Grade | yes
Functional Relationships |  |  |  |  | VAC | yes | Steam
Room Adjacencies | EM40 Labs |  |  |  | Spec'ty Gases |  | Type/Pressure
Room Activity | Ceramic Processing |  |  |  | Other | floor drains
Architectural |  |  |  |  | Electrical/Communication
Floor Material | Concrete |  |  |  | Note: Identify receptacle NEMA type, voltage, phase and load (HP, amps, voltamps or watts)
Base Material |  |  |  |  | Power | General | 110 v
Wall Material | Epoxy paint |  |  |  |  | Special | 240v (1/3 phase); 480v (3 phase)
Ceiling Material | Lay in | Height | 10'-0" |  |  | Isolated recept
Door Material | HM | Paint | Epoxy |  |  | Dedicated circuit | Yes | Describe: | each oven
Door Height | 7'-10" | Width | 6'-0" (2 leaf) |  |  | Spec'l Grounding | No | Describe:
Acoustics / Vibration | none |  |  |  |  | Emergency | No | Describe:
Floor Loading | na |  |  |  |  | UPS | No | Describe:
Other |  |  |  |  | Phone | yes
Lab Casework |  |  |  |  | Data | yes
Standing Ht. Bench | 90% | Sitting Ht. Bench |  | 10% | Safety
Island Upper | yes | Wall Upper |  | yes | Eyewash
Casework Material | metal | Countertop Material |  | epoxy | Emergency Shower
Tables | yes | Shelves |  |  | Fire Extinguisher |  | yes
Lighting |  |  |  |  | Fire Suppression System |  | sprinklers
Fixture Type | standard |  | Foot-candles | std | Fire Detection System |  | smoke detectors
Fixture Type | standard |  | Foot-candles | std | Gas Cylinders
Switching | standard |  |  |  | Type |  | Qty | Size
Mechanical | Standard |  |  |  | Type |  | Qty | Size
Temperature | standard | Range |  |  | Type |  | Qty | Size
Humidity | standard | Range |  |  | Type |  | Qty | Size
Cleanliness Class/Special Filtration |  | none |  |  | Chemicals
Pressurization (+/-) | neg | Range |  |  | Flammables
Air Changes/Hour | standard | Range |  |  | Acids
Emergency Pushbutton |  | Y/N |  |  | Bases
Fume Hood |  | Type |  |  | Comments
Point Exhaust |  | Equip Exhaust
Canopy Hood | 8 | Other
Equipment |  |  |  |  | Size |  | Electrical Req.
Name |  | Quantity | New/Existing | Bench/Floor/Wall | WxDxH | Weight | (See note above) | Piped Services | Vent
Carbolite Oven |  | 1 | existing | floor |  |  | 240v
Grieve Oven |  | 1 | existing | floor |  |  | 230v
Small Lindberg Oven |  | 1 | existing | floor |  |  | 240v
Cress Oven |  | 1 | existing | floor |  |  | 240v
Fisher Oven |  | 1 | existing | floor |  |  | 120v
Large Lindberg Oven |  | 1 | existing | floor |  |  | 230v
Sybron Oven |  | 1 | existing | floor |  |  | 240v
Medium Lindberg Oven |  | 1 | existing | floor |  |  | 220v

## 3.1 Nonmetal Mach
NASA 4602
Laboratory Design Criteria
Group | EM40 |  |  |  | Space Name |  | Nonmetals Machining Laboratory
Exist. Loc. | 4612/1123 & quarter 1109 |  | Date: | 2007-10-19 | Program Number
General |  |  |  |  | Plumbing
Unit Area | 530 |  |  |  | Sinks | deep epoxy | Acid Waste Sys | no
Qty. | 1 |  |  |  | Cold Water | yes | Hot Water | yes
Hrs Occupancy | 10 hrs/wk |  |  |  | Pure Water | no | Grade | i.e., type II or 18 Mgohm
# Occupants | 2 |  |  |  | Comp Air | yes | Pressure/Grade | 90 PSI minimum
Functional Relationships |  |  |  |  | VAC | yes | Steam | none
Room Adjacencies | EM40 Labs |  |  |  | Spec'ty Gases | no | Type/Pressure | no
Room Activity | machining of polymers, ceramics, and composites |  |  |  | Other | floor drain
Architectural |  |  |  |  | Electrical/Communication
Floor Material | painted concrete |  |  |  | Note: Identify receptacle NEMA type, voltage, phase and load (HP, amps, voltamps or watts)
Base Material |  |  |  |  | Power | General | 110 v at 6' on center
Wall Material | epoxy paint |  |  |  |  | Special | 230V, 230V, 600V
Ceiling Material | gyp bd./lay-in | Height | 10' |  |  | Isolated recept | yes
Door Material | hollow metal | Paint | Latex or Epoxy |  |  | Dedicated circuit | yes
Door Height | 7'-10" | Width | 6'-0" (2 leaf) |  |  | Spec'l Grounding |  | Describe:
Acoustics / Vibration | NOISY |  |  |  |  | Emergency |  | Describe:
Floor Loading | HEAVY |  |  |  |  | UPS |  | Describe:
Other |  |  |  |  | Phone | yes
Lab Casework |  |  |  |  | Data | yes
Standing Ht. Bench | 90% | Sitting Ht. Bench |  | 10% | Safety
Island Upper | no | Wall Upper |  | yes | Eyewash | yes
Casework Material | metal | Countertop Material |  | laminate/epoxy | Emergency Shower |  | no
Tables | yes | Shelves |  | yes | Fire Extinguisher
Lighting |  |  |  |  | Fire Suppression System |  | sprinklers
Fixture Type | fluorescent |  | Foot-candles | standard | Fire Detection System |  | smoke detectors
Fixture Type | tasklighting at lab benches |  | Foot-candles | standard | Gas Cylinders
Switching | Standard |  |  |  | Type |  | Qty | Size
Mechanical | Standard |  |  |  | Type |  | Qty | Size
Temperature | 72F | Range | +/- 2F |  | Type |  | Qty | Size
Humidity | controlled | Range | specify if reguired |  | Type |  | Qty | Size
Cleanliness Class/Special Filtration |  | Existing air filtration system (ceiling) |  |  | Chemicals
Pressurization (+/-) | neg | Range |  |  | Flammables | no
Air Changes/Hour | standard | Range |  |  | Acids | no
Emergency Pushbutton | yes if req'd | Y/N |  |  | Bases | no
Fume Hood |  | Type | Benchtop |  | Comments
Point Exhaust | yes | Equip Exhaust | yes |  | Need high power air conditioning to dissipate heat generated from machining carbon, etc.
Canopy Hood | yes | Other | Vac. Solid Waste system (exterior w/ ducting)
Equipment |  |  |  |  | Size |  | Electrical Req.
Name |  | Quantity | New/Existing | Bench/Floor/Wall | WxDxH | Weight | (See note above) | Piped Services | Vent
Monarch Lathe |  | 1 | E | Floor
Surface Grinder |  | 1 | E | Floor
Bridgeport Mill |  | 1 | E | Floor |  |  | 110v, 230v 3ph
Horizontal Band Saw |  | 1 | E | Floor |  |  | 230v
Romi Lathe |  | 1 | E | Floor |  |  | 600v 3ph
Bench Grinder |  | 1 | E | Floor |  |  | 110v
Belt Sander |  | 1 | E | Floor |  |  | 600v 3ph
Band Saw |  | 1 | E | Floor
Drill Press |  | 1 | E | Floor |  |  | 240v 30a
Overhead Lift |  | 1 | N | Ceiling

## 3.1 Elasto
NASA 4602
Laboratory Design Criteria
Group | EM 40 |  |  |  | Space Name |  | Elastomeric Materials Laboratory
Exist. Loc. | bldg. 4612 room 1125/1128 |  | Date: | 2007-03-26 | Program Number |  | 102
General |  |  |  |  | Plumbing
Unit Area | 990 |  |  |  | Sinks | 25x15x12 | Acid Waste Sys
Qty. | 1 |  |  |  | Cold Water | at sink | Hot Water | at sink
Hrs Occupancy | 10 |  |  |  | Pure Water |  | Grade
# Occupants |  |  |  |  | Comp Air | 90 psi shop | Pressure/Grade
Functional Relationships |  |  |  |  | VAC |  | Steam | To rubber mill
Room Adjacencies | EM40 Labs, Rubber Mill Areas |  |  |  | Spec'ty Gases |  | Type/Pressure
Room Activity | Wet Chemistry |  |  |  | Other | Floor drains. Piping interconnects between heating/cooling unit and mixer
Architectural |  |  |  |  | Electrical/Communication
Floor Material | Vinyl Tile |  |  |  | Note: Identify receptacle NEMA type, voltage, phase and load (HP, amps, voltamps or watts)
Base Material | Vinyle base |  |  |  | Power | General | NEMA 5-20, 120 volt, Single phse, 20 amps,
Wall Material | Epoxy paint |  |  |  |  | Special | provide lockable safety disconnects for all 480 volt equipment
Ceiling Material | Lay in | Height | 9'-6" |  |  | Isolated recept
Door Material | HM | Paint | Epoxy |  |  | Dedicated circuit | No | Describe:
Door Height | 7'-10" | Width | 6'-0" (2 leaf) |  |  | Spec'l Grounding | No | Describe:
Acoustics / Vibration | none |  |  |  |  | Emergency | No | Describe:
Floor Loading | na |  |  |  |  | UPS | No | Describe:
Other |  |  |  |  | Phone | yes
Lab Casework |  |  |  |  | Data | yes
Standing Ht. Bench | 90% | Sitting Ht. Bench |  | 10% | Safety
Island Upper | yes | Wall Upper |  | yes | Eyewash | potable water supplied eyewash, connected to drain
Casework Material | metal | Countertop Material |  | epoxy | Emergency Shower
Tables | yes | Shelves |  |  | Fire Extinguisher |  | As required by safety
Lighting |  |  |  |  | Fire Suppression System |  | sprinklers
Fixture Type | standard |  | Foot-candles | std | Fire Detection System |  | smoke detectors
Fixture Type |  |  | Foot-candles |  | Gas Cylinders
Switching |  |  |  |  | Type |  | Qty | Size
Mechanical | Standard |  |  |  | Type |  | Qty | Size
Temperature | standard | Range |  |  | Type |  | Qty | Size
Humidity | standard | Range |  |  | Type |  | Qty | Size
Cleanliness Class/Special Filtration |  | none |  |  | Chemicals
Pressurization (+/-) |  | Range |  |  | Flammables | need flammable cabinet
Air Changes/Hour |  | Range |  |  | Acids
Emergency Pushbutton |  | Y/N |  |  | Bases
Fume Hood |  | Type |  |  | Comments
Point Exhaust |  | Equip Exhaust |  |  | Need piping interconnects between mixer and heating / cooling units. Need chilled water supplied to heating / cooling unit at mixer. Need steam / hot water source at mill.
Canopy Hood |  | Other
Equipment |  |  |  |  | Size |  | Electrical Req.
Name |  | Quantity | New/Existing | Bench/Floor/Wall | WxDxH | Weight | (See note above) | Piped Services | Vent
Rubber Mixer |  | 1 | Existing | Floor | 94x45x99" | 1500 | from control cab | See comment | canopy hood, demand control
Rubber Calender Roll Mill |  | 1 | Existing | Floor | 90x56x84" | 1750 | 480volt, 50amp | See comment | canopy hood, damand control
Mooney Viscometer |  | 1 | Existing | Bench | 36X36x30" | 300 | 120volt, 20 amp
Weighing Cabinet |  | 1 | Existing | Floor | 48x48x60" | 175 | 120volt, 15 amp
Rubber Mixer circ. water heater |  | 1 | Existing | Floor | 17x40x26" | 100 | 480volt,20amp | See comment
Rubber Mixer circ.pump/cooler/reserv. |  | 1 | Existing | Floor | 28x36x34 | 250 | 480volt,15amp | Chilled water
Rubber Mixer power control cabinet |  | 1 | Existing | Wall | 30x16x42" | 100 | 480volt, 150amp | no
Portable HEPA dust collector |  | 1 | Existing | Floor | 20x30" | 40 | 120volt, 15 amp | no